## Slide 1: Testing scanning probes for PE THz EPR microscopy
(R. Hillenbrand – nanoGUNE)

## Slide 2: Outline
THz s-SNOM setup
THz magnetic probe
  General concept
  Simulations
  FIB fabrication
THz magnetic probe testing idea
First experimental step: THz electric field mapping
Conclusions

## Slide 3: Current THz scattering SNOM implementation
Source
THz Gas Laser (Coherent SIFIR 50)
Discret lines from 1.04 THz to 7 THz
Detector (currently implemented)
1 THz bandwidth at 2.5 THz
Room-temperature operation
1 mW optimal local oscillator power
Future detector (ordered)
Hot electron bolometer
Helium cooled (closed cycle)
Covers the range from < 1THz to 10 THz
Signal bandwidth > 250 kHz
Successfully tested with our setup
NeaSNOM
Commercial scattering SNOM
AFM based scanning technique with cantilevered tips (typicall tip oscillation frequencies 20 kHz-300 kHz)
Demodulation bandwidth ~1 MHz
Interferometer
beam splitter (polished silicon wafer)
Long travel delay stage
Provides homodyne signal amplification
Allows phase resolved detection
THz Laser
Mirror (on delay stage)
Tapping AFM

## Slide 4: Reasons for using tapping AFM
Einc
Ebg
Enf
Scattering is dominated by a strong background (tip, sample roughness, etc.): Esca = Ebg+Enf Ebg >> Enf
To extract small near-field signal:
  Higher harmonics demodulation
  Interferometric detection
Typical tapping amplitudes ~20-200nm incompatible with tuning fork

## Slide 5: Outline
THz s-SNOM setup
THz magnetic probe
  General concept
  Simulations
  FIB fabrication
THz magnetic probe testing idea
First experimental step: THz electric field mapping
Conclusions

## Slide 6: Concept: analogue to dielectric contract probing
“electric” hotspot
Einc
Esca
e
Detector signal can be used for mapping dielectric properties of materials
Esca= Esca(e)
Standard s-SNOM:
“Magnetic” s-SNOM:
“magnetic” hotspot
Einc
Esca
e,m
Detector signal can be used for mapping dielectric properties of materials
Esca= Esca(e,m)
Problems:
Need magnetic tip
No “magnetic” sample, i.e. magnetic transitions in THz/far-IR @ room-T

## Slide 7
apex
d
THz magnetic tip candidates
apex
tetrahedron tip
pyramid
cone with large radius
The distance between magnetic hotspot and the sample is limited by d ~ 200 nm, which significantly reduced the magnetic field at the sample compared to other designs
Harder to fabricate/mill
Good for numerical simulations
easier to fabricate based on commercial large-radii gold tips (radii 250 - 750 nm are available)

## Slide 8
Magnetic tip: Field maps
0
6000
H/H0
z=0
x
z
magnetic field (enhancement)
magnetic field enhancement along the yellow line
0
50
E/E0
z=0
electric field (enhancement)
electric field enhancement along the yellow
Si
Si
width for diabolo: w = 15 um
length for diabolo: L = 16.5 um
Height for Si tip: H = 15 um
thickness of gold: t = 25 nm
width for metal gap: wg= 50 nm
width for Si tip: wsi = 15 um
Illumination: l = 180 mm
z=0
z=0
Tip apex (bottom view)
q=50o
Einc
Hinc
30o
x
z
y
Top s-polarized illumination

## Slide 9
Magnetic tip: opening angle dependence (θ)
width for diabolo: w = 7 um
length for diabolo: L = 8 um
thickness of gold: t = 25 nm
width for metal gap: wg = 50 nm
Depth/height for Si tip changes as θ changes
Conclusion:
3D structure is better than the planar structure (cavity effect)
Large opening angles are preferred up to an optimum at about ? o
q close to the large-radius cone tip (40o)
Einc
Hinc
x
z
y
Bottom s-polarized illumination
q
Depth of Si tip

## Slide 10
Magnetic tip: Antenna length (L) depependence
Einc
Hinc
x
z
y
Bottom s-polarized illumination
q
Depth of Si tip
As expected:
The resonant wavelength red-shifts when L increases
Field enhancement increases when length L increases
open angle: θ = 50o
width for Si tip： wsi = 15 um
Height for Si tip: H = 15 um
thickness of gold: t = 25 nm
width for metal gap: wg= 50 nm
width of diabolo changes as L changes (L/w = constant)

## Slide 11: FIB fabrication: diablo antenna on commerical Si tip
Si
Au
Au
Au
Au
Si
Si
Si
Fabricated magnetic antenna tips using FIB patterning is relatively straightforward

## Slide 12: FIB rounded standard Si tip for large radius probes
Si
Si
Can fabricate blunt tips with large opening angle in house

## Slide 13: Commercial Au-coatedSi tips with large apex radius
Use FIB to mill the metal coating to obtain Diablo antenna

## Slide 14: Outline
THz s-SNOM setup
THz magnetic probe
  General concept
  Simulations
  FIB fabrication
THz magnetic probe testing idea
First experimental step: THz electric field mapping
Conclusions

## Slide 15: THz magnetic samples for testing THz magnetic tips
To test the “magnetic” tip performance in the absence of magnetic transitions at THz at room-T => use planar magnetic resonator structures:
  bar antennas
  diabolo antennas
  split rings
  complimentary structures?
These resonators can be excited via magnetic hot-spot of the tip:
i.e. magnetic field of the antenna
H
J
H
J
J
H

## Slide 16: Testing concept: cross-polarization measurements
Illumination: p-pol
Detection: s-pol
Illumination: s-pol
Detection: p-pol
Si
Au
Einc(p)
Low index substrate
Esca(s)
Si
Au
Einc(s)
Low index substrate
Esca(p)
Bottom illumination:
Si
Au
Einc(s)
Esca(s)
Not possible with our setup at the moment
Problem: Scattered THz electric field (which is detected in far field) could be affected by dielectric contrast and electric fields of the planar antennas

## Slide 17: Outline
THz s-SNOM setup
THz magnetic probe
  General concept
  Simulations
  FIB fabrication
THz magnetic probe testing idea
First experimental step: THz electric field mapping
Conclusions

## Slide 18: THz el. antenna near-field mapping
Au
Einc
Low index substrate
Esca
Si
Au
“Electric” tips
Mapping antenna fields at THz with s-SNOM is an uncharted territory.
Trying different tips concepts

## Slide 19: THz electric field mapping measurement results
Topography
Gold rod on Quarz (low index substrate)
Dimensions: 32 μm x 100 nm
Observations:
Mainly sensitive to out-of plane electric field above rod
Asymmetry
 We do not obtain the expected images. Reason unclear
Analogous experiment in mid-IR
Laser Photonics Rev., 1–13 (2015) /
DOI 10.1002/lpor.201500031
Experiment with THz s-SNOM
(with large radius FIB-milled silicon tip)
THz amplitude image
Concept:
Incident field (s-pol) excites electric dipole resonance in rod
Diel. tip scatters the electric antenna field
Tip-scattered field is detected
s-pol scattered field yields total antenna near field
p-pol scattered field yields vertical antenna near fields

## Slide 20: THz electric field mapping measurement results
|Ez+15|
Simulated field plotted 800 nm above sample
Asymmetry can be simulated by adding a constant value to the calculated out-of-plane electric field (indicates interference)
No indication that the in-plane electric field component (hotspot in the antenna gap) is mapped
Topography
Amplitude s2
Phase f2
|Ey+15|
Out-of-plane el. field amplitude (Ez)
In-plane el. field amplitude (Ey)
P-polarized illumination
Silicon tip
Bowtie antenna with 50 nm gap
Ein

## Slide 21: Conclusions/Problems/Outlook
s-SNOM required background suppression, which relies on relatively large tapping. Do we need background suppression for EPR?
We can work only with cantilevers, which could be mounted to tuning fork later. If background suppression based on tapping is required, tuning fork mounting needs to be adapted correspondingly.
Candidates for magnetic tip and antenna-based test sample are identified via simulations
THz is different from IR: preliminary el. antenna filed mapping results are unclear.
Simulations show large mag. enhancement factors for mag. tips, but the overall magnetic field could still be weak compared to the electric field. This could mask the magnetic response on the test samples with response due to dielectric permittivity and electric field components (mostly relevant for planar antenna sample).
Are there materials with magnetic transitions at THz/mid-IR and at room temperature to verify magnetic field at tip apex ?